## Slide 1
Can Internal Cooling Radiofrequency Catheters
with Local Hypothermia Eliminate
the Need for Tumescent Anesthesia?
: A Single-Center, Retrospective, Observational Study

## Slide 2
Minimally invasive procedures for treating chronic venous diseases have advanced significantly over the last two decades, resulting in reduced pain, less discomfort, and minimal interference with normal daily activities .
Radiofrequency ablation (RFA) is one of the most effective and safe endovenous treatment options for chronic venous disease. The ClosureFast system (Medtronic, USA) is the most commonly used RFA equipment and allows for the segmental ablation of veins.
However, the conventional system also has several limitations, one of which is the requirement for tumescent anesthesia. Tumescent anesthesia can prolong the procedure and may result in bruising and swelling of the lower limbs. Moreover, the local anesthetic agent used in tumescent anesthesia can sometimes carry the risk of systemic toxicity if the epinephrine component enters systemic circulation.
This study investigated the effectiveness and safety of tumescentless RFA using an internal cooling system with a VENISTAR catheter and local hypothermia.
Purpose

## Slide 3
This retrospective analysis enrolled 70 of 87 patients who had undergone endovenous RFA for above-knee (AK) great saphenous vein (GSV) between March 1, 2023 and November 30, 2023; 109 legs were evaluated.
The patients were categorized based on the use of tumescent anesthesia into two groups: Group T (conventional tumescent group, n =50) and Group N (non-tumescent group, n=59).
Group T received a tumescent anesthetic solution consisting of normal saline mixed with lidocaine, epinephrine, and sodium bicarbonate.
Group N underwent tumescentless RFA utilizing an internal cooling system and local hypothermia.
Postoperative pain (measured using the Visual Analog Scale), complications (including bruising, paresthesia, recanalization, revascularization, endothermal heat-induced thrombosis, and deep vein thrombosis), and operation times were assessed at 3 hours, 24 hours, and 10 days postoperatively.
Methods

## Slide 4
Methods
Figure 1. VENISTAR (STARmed, Korea) radiofrequency ablation (RFA) equipment. A) Illustration of VENISTAR equipment, radiofrequency generator, and catheter with internal cooling system (e.g., a circulating pump with 0°C saline).B) The RFA procedure using VENISTAR equipment. Internal cooling was achieved by using 0°C saline perfusate. The tip temperature of the catheter was kept low during ablation.
A)
B)

## Slide 5
Results
All endovenous RFA procedures were performed using a VENISTAR catheter targeting the GSV in the AK region. Due to the time requirements associated with tumescent anesthesia, the mean operative time in group T (42.3 minutes) was longer than that in group N (31.4 minutes) (p<0.01)
No instances of revascularization or recanalization of the treated vein were observed in either group 10 days post-operation. Postoperative outcomes did not differ significantly between the two groups.
There were no major complications related to endovenous RFA, such as DVT or EHIT, in either group. Postoperative bruising was observed in 6.8% (4/59) of patients in group N and 4.0% (2/50) of patients in group T (p=0.34). Additionally, 1.7% (1/59) of patients in group N and 6.0% (3/50) of patients in group T experienced postoperative paresthesia (p=0.27).
Postoperative pain scores, assessed with VAS scales at 3 h, 24 h, and 10 d post-operation, showed no difference between the two groups. Both groups reported only minor pain, with VAS scores consistently below 4.

## Slide 6
Other studies have indicated that tumescentless endovenous RFA with local hypothermia and compression is both safe and effective, with no associated adverse events, including thermal injuries.
We hope that this novel device can be utilized for the treatment of chronic venous disease, as it offers a shorter procedure time, is less invasive, and is technically more straightforward compared to conventional radiofrequency ablation (RFA) that uses tumescent anesthesia.
To the best of our knowledge, this is the first study to investigate the efficacy and safety of the internal cooling VENISTAR device in tumescentless RFA.
This innovative approach marks a recent advancement in endovenous RFA techniques, eliminating the need for additional tumescent procedures.
Conclusion